MelansonTech

Isolated Shunt Amplifier

Open-source isolated current measurement  ·  Project description, Rev 2

Document revised 20 August 2026 to match the Rev 2 schematic set and bill of materials. Specifications marked (V) are derived from the schematic and still need bench confirmation.

# Overview

The Isolated Shunt Amplifier is an open-source instrument front end that turns a high-side shunt voltage into a low-voltage signal your scope or DAQ can read safely, while keeping the measured circuit galvanically isolated from your test gear.

It is built around a Skyworks Si8921 isolated amplifier with a 10 mΩ four-terminal Kelvin sense resistor, a precision difference amplifier, and a relay-switched second gain stage. A slide switch selects off, the 10 A range (gain ×10), or the 5 A range (gain ×100); the board runs from a single 15 V supply and generates its own isolated 5 V rail on the measured side.

# Key features

- Two current ranges: 10 A at gain ×10 (1 V full scale) and 5 A at gain ×100 (5 V full scale), selected by a front-panel switch that also serves as the power switch.

- True galvanic isolation: the measured circuit is isolated from the output side by the Si8921 and an isolated DC-DC module, both 4.2 kV rated.

- Precision signal path: a Vishay WSL3637 10 mΩ, 3 W, four-terminal sense resistor, 0.05 % gain-setting resistors and a trimmable zero.

- Single supply: single 15 V input; the board makes its own +12 V, +5 V and isolated +5 V rails.

- Compact: 50 Ω SMA output, terminal-block current input, fits a small 3D-printed enclosure (panels included in the repository).

- Fully open: schematics, PCB, Gerbers, BOM, 3D models and simulation are all published.

# Specifications

| Parameter | Value | Notes |
| --- | --- | --- |
| 10 A range (gain ×10) | 10 A max | Switch position 2 — 1 V per 10 A, so 1.0 V at 10 A |
| 5 A range (gain ×100) | 5 A max | Switch position 3 — 1 V per 1 A, so 5.0 V at 5 A |
| Sense resistor | 10 mΩ, 3 W, ±0.5 %, ±50 ppm/°C | Vishay WSL3637R0100DEA, 4-terminal |
| Shunt voltage at max current | 100 mV (10 A) / 50 mV (5 A) | ±250 mV maximum at the isolator input |
| Isolator gain | 8.2 (fixed) | Si8921BD-IS4R |
| Total gain | ×10 or ×100 | Relay-switched second stage |
| Bandwidth, ×10 | ≈500 kHz (V) | Set by the Si8921, not by an on-board RC |
| Bandwidth, ×100 | ≈30–60 kHz (V) | Set by C8 and C9, switched with the range |
| Full-scale output | 1.0 V (10 A range) / 5.0 V (5 A range) | IsenseOut relative to ISenseRTN (chassis) |
| Maximum output | 5 V | Note: J1 is silkscreened ±4.5 V — see the design review |
| Output impedance | 50 Ω | R4 ∥ R8 — use a 1 MΩ input, not 50 Ω |
| Output connector | SMA female, edge launch | Cinch 142-0701-801 |
| Input common mode | 400 Vdc | Limited by the J2 terminal block rating |
| Isolation rating | 4.2 kV (Si8921 and PSU1) | Withstand rating, not continuous working voltage |
| Accuracy, 10 A range | ±(1.0 % of reading + 12 mA) (V) | At 23 °C ±5 °C. Typical ±(0.6 % + 9 mA) |
| Accuracy, 5 A range | ±(0.9 % of reading + 8 mA) (V) | At 23 °C ±5 °C. Typical ±(0.6 % + 5 mA) |
| Temperature coefficient | ±(0.011 % of reading + 0.5 mA)/°C (V) | Outside 23 °C ±5 °C |
| Operating temperature | 0 °C to 40 °C | Accuracy specified at 23 °C ±5 °C |
| Accuracy, user gain calibrated | ±(0.25 % + 12 mA) / ±(0.15 % + 8 mA) (V) | Needs a scale factor in your own system — there is no gain trim on the board |
| Auxiliary supply | 15 V, isolated, ≈300 mW | 2.1 × 5.5 mm barrel jack, centre positive |
| Signal ground reference | ISenseRTN, bonded to chassis | Board 0 V floats 6 V below chassis |
| Input connector | 2-pos 5.08 mm terminal block | Phoenix 5442002, 12 A / 400 V |

Required — the auxiliary supply must be isolated Use a 15 V supply whose output is floating: a two-pin double-insulated adapter, or a bench supply with genuinely isolated outputs. Do not use a supply whose negative output is bonded to earth. Use one isolated supply per board. Do not power two amplifiers from a single shared supply. Why: ISenseRTN is bonded to chassis and is the reference your instrument measures against, so the board's internal 0 V sits about 6 V below chassis. The auxiliary supply connects to that internal 0 V. If its negative is also earthed — as it is on USB-derived supplies, and on bench supplies with the negative terminal strapped to the earth post — the reference buffer is shorted to earth. The measurement collapses and the part overheats.

# Auxiliary supply requirements

The amplifier is powered from a single 15 V auxiliary input on the 2.1 × 5.5 mm barrel jack (centre positive) and draws roughly 300 mW. From it the board generates its own +12 V, +5 V and isolated +5 V rails.

The one hard requirement is isolation. ISenseRTN — the output return, on the SMA shell — is bonded to chassis, and everything the instrument reads is referenced to it. The board's internal 0 V rail therefore floats about 6 V below chassis, and the auxiliary supply's negative terminal connects to that floating rail. An auxiliary supply that bonds its negative output to earth ties the board's 0 V to chassis and shorts the reference buffer.

| Supply type | Suitable? |
| --- | --- |
| Two-pin double-insulated wall adapter | Yes — output is floating |
| Bench PSU with floating outputs | Yes — leave the negative terminal unstrapped |
| Bench PSU with negative strapped to earth | No — remove the strap first |
| USB-derived or PoE-derived 15 V | No — host ground is normally earthed |
| One supply shared between two boards | No — use one isolated supply per board |

A brief note on the reasoning behind the arrangement: bonding the 6 V reference to chassis is what lets a single positive rail produce a genuinely ground-referenced ±4.5 V output, with no negative regulator or charge pump. The isolated auxiliary supply is the price of that simplicity.

Two further practical points. The output source impedance is 50 Ω (R4 ∥ R8), so an instrument that terminates into 50 Ω will read half the true output — use a high-impedance (1 MΩ) input. And the zero is trimmed with R28, which adjusts DiffAmpRef against the fixed chassis reference; trim it after the board has reached thermal equilibrium, since the residual drift after warm-up is what limits the zero stability.

# Accuracy

The switch legends are scale factors, not limits. The ×10 range runs to 10 A for a 1.0 V full-scale output; the ×100 range runs to 5 A for a 5.0 V output, which is where the amplifier's 5 V ceiling stops it. The figures below are percentages of each range's own full scale.

Before the zero trim is set, the static offsets in the chain add up to 2.9 % of full scale on the 10 A range and 4.5 % on the 5 A range. Trimmer R28 nulls every one of them in a single adjustment, which is why it is a functional part of the design rather than a refinement.

| Total error, typical, as shipped | 10 A range (×10) | 5 A range (×100) |
| --- | --- | --- |
| At full scale | ±69 mA (0.69 % of reading) | ±36 mA (0.73 %) |
| At 50 % of range | ±40 mA (0.79 %) | ±22 mA (0.88 %) |
| At 10 % of range | ±16 mA (1.62 %) | ±10 mA (2.09 %) |

The specification is written as ±(% of reading + fixed mA) because the two parts behave differently: the proportional part scales with the reading, while the fixed part does not. The fixed part is not un-nulled offset: R28 zeroes the static offset completely, and what remains is only what zeroing cannot reach — drift that appears after the trim, Si8921 nonlinearity, which is not an offset at all, and the finite resolution of the pot. Below about a third of range the fixed part dominates and the error as a percentage of reading climbs past 1 % — normal for an instrument of this class, and the reason a single percentage would flatter it.

The proportional term is not a linear worst-case sum. A Monte Carlo over two million boards, with every tolerance uniform within its limit, never once reached that sum — the median was 22 % of it and the 99.9th percentile 73 %. But pure RSS is too optimistic, because boards built from one reel of shunts share the same offset from nominal. The published figure holds the two batch-correlated terms — the shunt's ±0.5 % and the isolator's ±0.2 % gain error — at their limits and combines everything else in quadrature.

The fixed term assumes the board has moved ±20 °C from the temperature at which R28 was trimmed. Re-zeroing at the working temperature roughly halves it, to ±8.6 mA and ±4.6 mA, at which point it is almost entirely Si8921 nonlinearity and no further trimming helps. Worth doing before a careful measurement.

Two observations from the budget. The shunt's ±0.5 % initial tolerance and the isolator's ±0.2 % gain error together dominate the worst case. Both are systematic per unit, but there is no gain trim on the board, so they remain in the shipped instrument — measuring one known current and applying a scale factor in your own system removes both and roughly triples the accuracy. And the two ranges come out within a few hundredths of a percent of each other: the ×100 range gives away twice the offset weight but takes back four times less shunt self-heating, and the two roughly cancel. The full working model is in the tolerance budget spreadsheet alongside this document.

# Intended use

The amplifier is aimed at bench and integration work where isolation, precision and a usable bandwidth all matter at once:

- High-voltage power systems: measure shunt current on a rail sitting hundreds of volts above your instrument ground.

- Electric vehicle battery testing: pack-level and module-level current measurement with a safe, isolated output.

- Industrial automation: isolated signal conditioning for data acquisition in equipment where the measurement point is not at ground.

- General laboratory use: a low-cost isolated current probe for lab work, with a 10 A and a 1 A range on one board.

# Repository contents

| Path | Contents |
| --- | --- |
| *.PrjPcb | CircuitStudio project files, Rev 1 and Rev 2 |
| *.SchDoc | Schematic sheets — see the README for which sheet belongs to which revision |
| *.CSPcbDoc | PCB layout (both projects build the same board file) |
| Default Configuration/Outputs/ | Gerbers, NC drill, BOM, pick-and-place, STEP export, schematic PDF |
| 3d Models/ | STL and STEP models — board, enclosure panels, bumpers |
| Simulation/ | LTspice model of the amplifier chain |
| docs/ | This document, DESIGN_REVIEW.md, and the tolerance budget spreadsheet |
| archive/ | Superseded sheets and layouts, referenced by neither project |

# Open source

This project is licensed under the CERN Open Hardware Licence Version 2 — Strongly Reciprocal (CERN-OHL-S-2.0). Copyright © 2024–2026 MelansonTech (Shawn Melanson).

This source describes Open Hardware and is licensed under CERN-OHL-S v2. You may redistribute and modify it under the terms of that licence. This source is distributed WITHOUT ANY EXPRESS OR IMPLIED WARRANTY, INCLUDING OF MERCHANTABILITY, SATISFACTORY QUALITY AND FITNESS FOR A PARTICULAR PURPOSE. Please see the licence for applicable conditions.

In practice: you may build, modify and sell this design, but a modified version — or a product made from one — must be published under the same licence. Improvements to the layout, alternate shunt options, enclosure variants, and build reports from your own bench are all welcome. Open an issue or a pull request.

Design work is done in Altium CircuitStudio. The files under Default Configuration/Outputs are generated artifacts; regenerate them from the project rather than editing them by hand.

# Known issues in Rev 2

These are documented in full in DESIGN_REVIEW.md and are open for Rev 3:

- the isolated auxiliary supply requirement is not marked on the board or the enclosure — only in this document.

- with a 15 V supply, the 78L12 regulator sits close to its dropout threshold once the input Schottky drop is accounted for.

- J1 is silkscreened ±4.5 V, but the 5 A range reaches 5.0 V at full scale — the marking and the range limit need reconciling.

- there is no gain trim on the board, so shunt and isolator initial tolerance stay in the shipped accuracy figure.

- the 100 kΩ reference dividers give the LM7321 buffers a high source impedance, so bias-current drift after warm-up limits how well the trimmed zero holds. 10 kΩ dividers would reduce this tenfold.

- R28 is annotated as 100 Ω on the schematic but specified as a 10 kΩ trimmer in the BOM; the BOM value is correct.

- the published 500 Vdc common-mode figure exceeds the 400 V rating of the input terminal block. This document uses 400 Vdc.

- C8, C43 and R23 are generic placeholder lines in the BOM and need real part numbers before a production run.

Designed by Shawn Melanson  ·  MelansonTech  ·  Rev 2 documentation, August 2026