## 1 Board
Test record — MelansonTech Isolated Shunt Amplifier Rev 2
One copy per board. Fill the yellow cells; everything else calculates. See docs/CALIBRATION.md for the procedures.
Board serial / identifier |  | Whatever you mark on the board
Date tested |  | YYYY-MM-DD
Tested by
Assembly revision | Rev 2
AMBIENT AND CONDITIONS
Ambient temperature at test |  | °C — the spec is written at 23 °C ±5 °C
Warm-up before trimming (minutes) |  | 30 minutes minimum, see CALIBRATION.md
Auxiliary supply voltage, measured |  | V — measure it at the jack. 13–15 V works; 15 V is needed for the full spec
Auxiliary supply isolated? |  | MUST be floating. See SAFETY.md and the README
EQUIPMENT USED
DMM / scope model and serial
Current source or reference ammeter
Reference ammeter accuracy |  | Needs to be better than the board to mean anything
Instrument input impedance | 1 MΩ | A 50 Ω terminated input reads HALF

## 2 Bring-up
Bring-up — first power-on
J2 disconnected, isolated supply, current limit ~200 mA. Order matters: work down the sheet.
Step | Check | Expected | Measured | Result
1 | Visual — solder bridges at J1, J2, J4 | none
2 | Visual — U2 and PSU1 orientation | correct
3 | SW1 position 1 (Off): supply current | ≈ 0 mA
4 | SW1 position 2 (10 A range): supply current | 90–100 mA
5 | SW1 position 2: D5 lit, D4 dark | exactly one lit
6 | SW1 position 3 (5 A range): D4 lit, D5 dark | exactly one lit
7 | SW1 position 3: Rel1 audibly operates | click heard/felt
8 | Nothing overheating (U6 warm is normal) | no fast heating
9 | Output near zero before trim | within ±50 mV
RAIL CHECK
| VregIn | ≈ Vin − 0.37 V
| +12 | 12 V if Vin ≥ 14.9 V, else ≈ Vin − 2.07 V
| +5 | 5 V ±5 %
| +5_ISO | 5 V, isolated side
| ISenseRTN, relative to board 0 V | ≈ 6 V
| DiffAmpRef, relative to board 0 V | ≈ 6 V
The +12 check reads your measured supply from the Board sheet. Above 14.9 V it expects a regulated 12 V; below that it expects an unregulated rail at roughly Vin − 2.07 V, which is correct behaviour, not a fault. See finding 2 in DESIGN_REVIEW.md.

## 3 Zero trim
Zero trim — R28
J2 open or shorted, 30 minutes warm-up, SW1 position 3 (5 A range, ×100). A null on this range holds on the other.
Temperature at which the trim was made |  | °C | Record it — the residual drift from here is what limits zero stability
Output before trimming |  | mV | IsenseOut relative to ISenseRTN
Output after trimming |  | mV | As close to zero as you can hold
R28 value fitted | 100 | Ω | Should be 100 Ω. The BOM's 10 kΩ line is an error
Trim position within travel |  |  | centre / near one end / at the end
Residual within ±5 mV? |  |  | A few mV is normal. If it will not null at either extreme of travel, fit a 200 Ω trimmer — see CALIBRATION.md.
If the trim runs out of range: the chain's static offset exceeds ±3 mV. This is a tolerance
stack, not a fault — at datasheet limits it can reach ±24 mV. Fit 200 Ω, which doubles the
range and still resolves to 0.12 % of full scale. Note it in the comments so the population
is tracked.
Re-trim whenever the board has moved temperature and small readings matter. Re-trimming at
the working temperature takes the fixed error term from 11.3 mA to 9.1 mA on the 10 A range.

## 4 Accuracy 10A
Accuracy — 10 A range (×10)
SW1 to this range. Nominal 100 mV/A, full scale 10 A. Published spec: ±(1.0 % of reading + 12 mA). Zero trim must be done first.
Point | True current (A) | Output measured (V) | Indicated current (A) | Error (mA) | Error (% of reading) | Spec limit (± mA) | Result
1 | 0.5 |  |  |  |  | 17
2 | 1 |  |  |  |  | 22
3 | 2 |  |  |  |  | 32
4 | 5 |  |  |  |  | 62
5 | 7.5 |  |  |  |  | 87
6 | 10 |  |  |  |  | 112
Overall
GAIN CONSTANT (optional — takes the proportional term to ±0.25 % / ±0.15 %)
k = true / indicated, from the highest point measured
Expect k within about ±1 % of 1.000. If k ≈ 2.0 your instrument is terminating into 50 Ω — use a 1 MΩ input.

## 5 Accuracy 5A
Accuracy — 5 A range (×100)
SW1 to this range. Nominal 1000 mV/A, full scale 5 A. Published spec: ±(0.9 % of reading + 8 mA). Zero trim must be done first.
Point | True current (A) | Output measured (V) | Indicated current (A) | Error (mA) | Error (% of reading) | Spec limit (± mA) | Result
1 | 0.25 |  |  |  |  | 10.25
2 | 0.5 |  |  |  |  | 12.5
3 | 1 |  |  |  |  | 17
4 | 2 |  |  |  |  | 26
5 | 3.5 |  |  |  |  | 39.5
6 | 5 |  |  |  |  | 53
Overall
GAIN CONSTANT (optional — takes the proportional term to ±0.25 % / ±0.15 %)
k = true / indicated, from the highest point measured
Expect k within about ±1 % of 1.000. If k ≈ 2.0 your instrument is terminating into 50 Ω — use a 1 MΩ input.

## 6 Optional
Optional characterisation
None of this is needed to release a board. All of it is currently UNMEASURED for the design as a whole — results here would close real gaps in the documentation.
BANDWIDTH — design review finding 8 says the ×100 figure is not obviously set by any RC
−3 dB bandwidth, 10 A range (×10) | kHz | Documentation claims ≈500 kHz, set by the Si8921
−3 dB bandwidth, 5 A range (×100) | kHz | Documentation claims ≈30 kHz; the schematic suggests 30–60 kHz
DRIFT — the whole fixed error term rests on this
Zero at trim temperature | mV | Baseline
Zero after +15 °C soak | mV | Predicted ≈ 4.6 mV shift at the output, 5 A range
Zero after −15 °C soak | mV
DiffAmpRef − ISenseRTN at trim temp | mV | Measuring this over temperature replaces an ESTIMATE in the budget
DiffAmpRef − ISenseRTN after +15 °C | mV | Directly measures the reference-drift term
THERMAL — design review finding 5
U6 case temperature, 30 min, still air | °C | Predicted ~0.6 W dissipation in a SOT-89
R5 shunt temperature at 10 A | °C | Predicted ~33 °C rise; drives the self-heating error term
Ambient during thermal measurements | °C
SUPPLY MARGIN — design review finding 2
+12 rail at 15.0 V in | V
+12 rail at 14.0 V in | V | If this sags, U1 is in dropout — the whole point of finding 2
Lowest input voltage holding +12 V | V | The real dropout threshold for this build
ISOLATION — see SAFETY.md before attempting anything here
Measured barrier creepage, closest point | mm | From the layout or the bare board. NOT yet verified for this design
Measured barrier clearance, closest point | mm
Hipot / isolation test performed? |  | Only with proper equipment and precautions

## 7 Summary
Summary and sign-off
Pulls the result from each sheet. Nothing to fill in except the comments and the sign-off.
Board serial | 0
Date | 0
Ambient at test (°C) | 0
Bring-up | not done
Zero trim residual (mV) | 0
Zero trim
Accuracy, 10 A range
Accuracy, 5 A range
Gain constant k, 10 A range
Gain constant k, 5 A range
OVERALL
Board result | INCOMPLETE
Comments / anomalies
Tested by
Date
Reviewed by
Write the two gain constants on the board or its label. A calibration nobody can find is a calibration nobody uses.